# Lecture 11 – Sensory neural prosthesis and transcutaneous energy

FIRST COCHLEAR IMPLANT (1978)

- Electrode array goes into cochlear

- Made from copper coil = inductor

  - Encapsulated in epoxy (may not have been medical grade)

  - Altogether went into electrode array

- Understanding risks:

  - By today’s standards, device would not be commercially marketed/implanted

  - Back in 1978, had a huge demand → advocated for use

  - 2/3 prototypes failed, but since one worked  → Cochlear still gained approx. 65% of world market

    - Implanted in > 350,000 people

- Reference to picture (right)

  - Cochlear = bony material

    - Size of tip of pinky

  - Need some material that fits in, loops around and draws out electrical stimulation from electrodes

  - Target tissue = yellow = spiral ganglion cells

    - Collect and form auditory nerve

    - Goes to sound centers of brain to be interpreted

  - Action potential

    - Needed to transport signals

ORIGINS OF NEUROMODULATION

- Manipulate cells’ behavior by changing the extracellular potential

- Eg. Applying a voltage relative to the inside of a cell = membrane becomes charged

  - Action potential  → enables certain functions in body to occur

SENSORINEURAL HEARING LOSS

- Sound hits the tympanic membrane  → vibration

  - Ossicle augments vibration  → creates greater mechanical movement on middle ear

  - Middle ear pushes onto cochlea, moving fluid around  → fluid movement bends into inner hair cells  → bending generates action potential ∴ sending signals to the brain

  - OCCURS IN THE ORGAN OF CORTI

- Target neurons for implant = spiral ganglion cells

  - Attached to hair cells

  - In diseases (eg. Meningitis) where hair cells are killed  → it cannot cause action potential from hair bending

- IF auditory nerve is present and functional, it can be electrically stimulated  → can detect sound

- Cochlear implant system

  - Outside = microphone + headpiece

  - Implantable cochlear stimulator  → cable  → through round window, into cochlear  → electrically stimulate various positions of cochlear

THE BIONIC EAR

- Stats:

  - Approx. 120 million are hearing impaired (Approx 1 million in Aus.)

  - 2% are fully deaf  → (2.4M worldwide, approx. 20K in Aus.)

  - FDA, initially let Cochlear implant into those profound deaf, as they thought it did more damage than good

THE STIMULUS WAVEFORM

- Use constant current pulses

  - Other devices typically use voltage

- Phase 1 does all the work

  - Gets action potential firing by injecting charge (coulombs)

- Inter-phase gap

  - Can be adjusted to what is most effective

  - Needed to allow time for action potentials to propagate down to designated area

- Phase 2

  - Charge recovery = reverse the induced charge and restore neutrality → ear is stimulated again without difficulties arising in electrochemistry

  - Without it, cause buildup of random chemicals and unwanted substances

  - Ideally, has the same amount of charge in second phase

    - But, in reality, not exactly equal due to ions/particles moving away from electrode site

  - Can still cause electrical stimulations

    - Solution: reduce current amplitude but have a pulse occur over a longer amount of time (area = 0)

- Charge = area under curve (sum must = 0, to be neutral)

- Inter-stimulus gap

  - Separation between pulses

  - Limited to how quickly you can feed information, and how quickly the electrodes are serviced

- Stimulation can only occur once at a time

  - Eg. If you have 22 electrodes, only one can be stimulated at a time

NUCLEUS CI24RE

- 24 channels = 22 channels on electrode array + little monopolar reference electrode (ball + plate)

- Monopolar ref. electrodes = return path for the negative return of electrical stimulation

- Contact materials = silicone, titanium and platinum

- All electronics are placed in a little capsule = hermitic chamber

  - Connects to each of little wires

- Sound processor has 2x microphones

  - Picks up surround sound

  - Assumes most important sound is in front,  so sound received = front minus behind sound

  - Sound processor → sends instruction and power to the RF link

- RF link

  - Via inductive coupling

  - Outside TX (transmitting) coil → RF signals go through skin → implant (circuitry grabs energy off) → detects when energy is turned on/off

  - Use RF coil to prevent infection from implanting

  - Don’t implant battery (due to big size); coils are very small → allowing babies to wear them

ANATOMY OF COCHLEA

- Scala media (media = middle)

- Scala vestibula

- Scala tympani

  - Tympani membrane is where hair cells are excited

  - Target area = spiral ganglion cells

- In picture [right]; electrode array (grey) = metal

  - If e-pulses are emitted from metal, we can excite spiral ganglion cells and send signals off to the brain to be interpreted as sound

- Electrode array

  - Used to be straight, but presently curled

  - Cochlea is curled

  - Allows electrode to get close as possible to spiral ganglion cells

  - Electrical stimuli is directed inside curl of electrode; when it was straight, it will go outside

- BUT: looking at x-ray image, there is still so much of cochlea that isn’t reached

  - Consequences:

    - Only detecting high pitch/frequency (low frequency = long wavelength penetrating through the deeper end of the cochlea

  - Goal:

    - Have widely spaced electrodes throughout the cochlea

    - Hear all sounds of the spectrum

  - Electrodes are limited to a range of frequencies; not a 1:1 correspondence (eg. Piano keys slammed hard)

  - In 25 years, < 7% of devices have failed